Berkshire & Surrey Pathology Services

Frimley Park Hospital

Portsmouth Road

Frimley GU16 7UJ

3rd August 2022

Dear Colleagues

RE Implementation of new analytical equipment: Go Live 22nd August 2022

Further to our previous correspondence, we are delighted to inform you that we have agreed a date for the switch over to our new Abbott analytical equipment at Frimley Park Hospital. This will further increase the quality of our service offered and to enhance our service resilience.

The change of equipment will be associated with a change in results, reference ranges and interpretation of a number of analytes. Analytical results measured on one piece of equipment do not necessarily give the same values when measured on a second, differing piece of equipment.

Following feedback we have included the average observed bias for the most significantly affected analytes. Any analytes not referred to are affected by <5% relative to the previous method.

| Analytes affected by between 5-10% |  |  |
| --- | --- | --- |
| ALT | Lactate | Testosterone |
| AST | Paracetamol | Progesterone |
| CK | Bilirubin | NT-pro BNP |
| Conjugated Bilirubin | Triglyceride | TSH |
| Oestradiol | LH | Digoxin |

| Analytes affected by >10% |  |  |
| --- | --- | --- |
| Albumin (-17%) | *LDH (-46%) | PTH (21%) |
| GGT (14%) | Lipase (-25%) | *SHBG (14%) |
| *Cortisol (16%) | DHEAS (14%) | Vitamin B12 (-17%) |
| Ferritin (34%) | *fT3 (-17%) | Carbamazepine (-13%) |
| *fT4 (-13%) | FSH (-20%) | Amylase (-11%) |
| Procalcitonin (17%) | *Prolactin (66%) | Vancomycin (-15%) |

For those assays denoted with ‘*’ more detailed information regarding assay change is available upon request.

The results for the following analytes are not directly comparable between different equipment and there is no clear bias observed between the old and new method:

Not directly comparable:

hs-Troponin I *

CA199 *

CA125 *

CA153 *

CEA *

Note AFP/PSA have been removed from the above list since previous correspondence due to additional in-house data suggesting excellent correlation

Reference range changes

Details of the reference range changes associated with this change can be provided upon request

Paediatric reference ranges

The implementation of the Abbott Alinity analysers will enable the laboratory to adopt the Internationally-recognised and evidence-based Caliper reference ranges for a number of analytes. A summary of these can be provided upon request

We can assure you that the appropriate reference range will be appended to all results regardless of the site of analysis.

We understand that these changes in results and associated reference ranges will create challenges in the ongoing monitoring of chronic conditions at and around the time of the implementation, and caution will be necessary in the interpretation of these results.

If you have any specific questions regarding the implementation of the equipment or any specific patient pathways that you feel may be impacted, please do not hesitate to contact the department by emailing Lisa Glendenning our Comms and Change Lead for the Automation Programme lisa.glendenning1@nhs.net.

Kind Regards,

Patricia Ruwona

BSPS Chief Operating Officer

and

David Cartwright

BSPS Associate Medical Director (Blood Sciences)