Berkshire & Surrey Pathology Services

Frimley Park Hospital

Portsmouth Road

Frimley GU16 7UJ

5th August 2022

Dear Colleagues

Abbott Alinity assay biases – Tumour marker assays

Summary

AFP results will be comparable with previous, with new age-related reference ranges for paediatrics.

CEA results will be on average 40% higher with a new reference range.

CA125 results will be on average 35% higher. This may impact ultrasound referrals.

CA153 results will be on average 15% higher.

CA199 results will not be comparable with previous, and there will be a new reference range.

PSA results will be comparable with previous, with new age-related reference ranges based on the NICE NG12 guidance1.

For all of these analytes, unpredictable biases may be seen for individual patients and re-baselining of some patients who are being monitored will be carried out by the laboratory.

Introduction

In February 2022, Berkshire and Surrey Pathology Services (BSPS) signed a contract with Abbott diagnostics to enable the acquisition and installation of new automation across the Pathology network. This will result in significant improvement to the laboratory’s robustness and resilience over the forthcoming years. The entire BSPS Network will be rolling out the Abbott Alinity suite of analysers over a period of 12 months across all sites.

One aspect of installing new automation is the necessity of moving to alternative reference ranges and methodologies as dictated by the manufacturer and evidence-based studies performed on the new instrumentation. There are several key assays where results will be significantly different and clinical teams will need to become very familiar with the new cut-offs in place.

Abbott Alinity

The majority of the diagnostic accuracy and clinical data for this assay has been generated using Abbott’s Architect platform. Abbott’s subsequent analytical platform the Alinity series of analysers use identical reagent and calibrants to the Architect and studies have found that results on both platforms are equivalent2.

Expected changes in results

AFP results show good correlation with the old method, with in-house studies suggesting an average bias of <10% for the vast majority of patients.

CEA results will be on average 40% higher

CA125 results will be on average 35% higher, but at very high levels results are higher

CA153 results will be on average 15% higher

CA199 results will not be comparable with the previous method, with variable biases between individuals.

PSA results show good correlation with the old method, with in-house studies suggesting an average bias of <10% for the vast majority of patients.

Impact on monitoring of patients

After the new methods are introduced, the laboratory will identify patients who are being monitored where there has been a change between the previous and current result which is not in line with that expected as detailed above. These samples will be analysed on both the old and new method, and both results will be reported to allow re-baselining of these patients. We will continue this for several months after the new methods are introduced. For CA199, all patients being monitored will be re-baselined.

Please contact the Duty Biochemist (fhft.bspsdutybiochemistteam@nhs.net) if there are any concerns around unexpected changes in tumour marker levels during the period after introduction of the new method.

Impact on suspected ovarian cancer referrals

CA125 results are expected to increase by approximately 35%, including around the cut-off of 35 kU/L given in NICE NG12. This may result in an increase in ultrasound referrals.

Changes to reference ranges

The following new reference ranges will be introduced for AFP, CEA, CA199 and PSA.

|  | AFP reference range (kU/L) |
| --- | --- |
| <1m | 0-1653 |
| 1m-<6m | 8-1123 |
| 6m-<1y | 0-85 |
| 1-<19y | 0-12 |
| Adult (≥19y) | 0-9 |

|  | CEA reference range (ug/L) |
| --- | --- |
| Non-smokers | 0-4 |

|  | CA199 reference range (kU/L) |
| --- | --- |
| All | 0-36 |

Reference ranges for PSA will change to align with those recommended in the NICE NG12 guidance1.

|  |  | PSA reference range (ug/L) |
| --- | --- | --- |
| Male | <40y | Use clinical judgement (NG12) |
|  | 40-49y | 0-2.50 |
|  | 50-59y | 0-3.50 |
|  | 60-69y | 0-4.50 |
|  | 70-79y | 0-6.50 |
|  | >79y | Use clinical judgement (NG12) |
| Female |  | None |

References:

- NICE guideline [NG12] Suspected cancer: recognition and referral. Last updated 15 December 2021

- Jong Do Seo, Da Young Song, Youngwon Nam et al, Evaluation of analytical performance of Alinity i system on 31 measurands, Practical Laboratory Medicine, Volume 22, 2020, https://doi.org/10.1016/j.plabm.2020.e00185

If you have any specific questions regarding the implementation of the equipment or any specific patient pathways that you feel may be impacted, please do not hesitate to contact the department by emailing Lisa Glendenning our Comms and Change Lead for the Automation Programme lisa.glendenning1@nhs.net.

Kind Regards,

Patricia Ruwona

BSPS Chief Operating Officer

and

David Cartwright

BSPS Associate Medical Director (Blood Sciences)