Issue: False positive results for Vero cytotoxin producing  Escherichia coli (VTEC)

Recipients: All Frimley ICB GP Practices

We are writing to inform you that we have identified and investigated an issue with the assay used to detect gastrointestinal pathogens by PCR of stool samples, which has produced false positive results for Vero cytotoxin producing  Escherichia coli (VTEC). The results for other pathogens in the panel are unaffected.

The root cause of this fault has been identified as a problem with the reagents supplied from the manufacturer.

A look-back exercise has been conducted, which has concluded that any stool results reported as VTEC DECTECTED by the Microbiology laboratory at Frimley Park Hospital from 01/09/2023 and by the Microbiology laboratory at Wexham Park Hospital from 01/10/2023 may be false-positives.

We have worked closely with the manufacturer to understand and rectify the issue.

Please be aware that any VTEC DETECTED results issued from today will be reported with the following comment until we have full assurance that the problem has been rectified: "Interpret result with caution due to ongoing issue. Await reference laboratory confirmation".

Please consider reviewing patients with non-resolving diarrhoea or abdominal symptoms where a VTEC DETECTED result has been reported during the periods of concern, in case further investigation may be indicated to exclude other causes. If a clinical conversation would be helpful, please contact your local duty medical microbiologist.